## Corrected Stroke
Bin | Frequency
4 | 0
4.2 | 0
4.4 | 0
4.6 | 0
4.8 | 0
5 | 63
5.2 | 281
5.4 | 34
5.6 | 0
5.8 | 0
6 | 0
More | 0

## Parallel Resistance
Bin | Frequency
0 | 0
2000000000 | 4
4000000000 | 356
6000000000 | 10
8000000000 | 7
10000000000 | 1
12000000000 | 0
14000000000 | 0
16000000000 | 0
More | 0

## Series Resistance
Bin | Frequency
15 | 0
15.5 | 0
16 | 0
16.5 | 35
17 | 165
17.5 | 139
18 | 39
18.5 | 0
19 | 0
More | 0

## HP Cap.
Bin | Frequency
0.6 | 0
0.8 | 0
1 | 20
1.2 | 358
1.4 | 0
1.6 | 0
More | 0

## Fluke Cap.
Bin | Frequency
0.8 | 0
1 | 0
1.2 | 0
1.4 | 378
1.6 | 0
1.8 | 0
More | 0

## SN357-01
357CH NJIT DM #357-001 5450-211 | [7 empty] | Last Updated: 1/29/10 |  | Temperature Sensitivity
Post Discrete Burn-in for XIRE0313MS6 | [9 empty] | %/C | 2.6%
Blocks: 1087 | [5 empty] | QC Tested
Serial # | Block # | Actuator Size | Manual Setup | Automated Setup | Test Temp (C) | Stroke @ 90V (microns) | Cap. Fluke (μF) | Cap @ 1Khz | Series Resistance (ohms) | Parallel Resistance (ohms) 10 sec 90V | Stroke at 90V & 24C (microns) | Relative Humidity | Comments | Count
1 | 1087 | XIRE0313MS6 |  | x | 21.2 | 5.4 | 1.34 | 1.01 | 17.43 | 6720000000 | 5.0069 | 14 |  | 1
2 | 1087 | XIRE0313MS6 |  | x | 21.2 | 5.4 | 1.34 | 1.01 | 17.91 | 2170000000 | 5.0069 | 14 |  | 2
3 | 1087 | XIRE0313MS6 |  | x | 21.2 | 5.4 | 1.33 | 1.01 | 17.77 | 7460000000 | 5.0069 | 14 |  | 3
4 | 1087 | XIRE0313MS6 |  | x | 21.2 | 5.6 | 1.33 | 1.02 | 17.26 | 2750000000 | 5.1923 | 14 |  | 4
5 | 1087 | XIRE0313MS6 |  | x | 21.2 | 5.4 | 1.35 | 1.02 | 17.64 | 2390000000 | 5.0069 | 14 |  | 5
6 | 1087 | XIRE0313MS6 |  | x | 21.2 | 5.4 | 1.35 | 1.02 | 17.72 | 2380000000 | 5.0069 | 14 |  | 6
7 | 1087 | XIRE0313MS6 |  | x | 21.2 | 5.4 | 1.33 | 1 | 17.86 | 2650000000 | 5.0069 | 14 |  | 7
9 | 1087 | XIRE0313MS6 |  | x | 21.2 | 5.5 | 1.36 | 1.02 | 17.54 | 2450000000 | 5.0996 | 14 |  | 8
10 | 1087 | XIRE0313MS6 |  | x | 21.2 | 5.4 | 1.31 | 1 | 17.59 | 2570000000 | 5.0069 | 14 |  | 9
11 | 1087 | XIRE0313MS6 |  | x | 21.2 | 5.5 | 1.36 | 1.03 | 17.51 | 2300000000 | 5.0996 | 14 |  | 10
12 | 1087 | XIRE0313MS6 |  | x | 21.2 | 5.5 | 1.31 | 1.01 | 17.35 | 2520000000 | 5.0996 | 14 |  | 11
15 | 1087 | XIRE0313MS6 |  | x | 21.2 | 5.4 | 1.35 | 1.03 | 17.28 | 2990000000 | 5.0069 | 14 |  | 12
16 | 1087 | XIRE0313MS6 |  | x | 21.2 | 5.5 | 1.35 | 1.03 | 17.42 | 2310000000 | 5.0996 | 14 |  | 13
17 | 1087 | XIRE0313MS6 |  | x | 21.2 | 5.4 | 1.35 | 1.02 | 17.51 | 2400000000 | 5.0069 | 14 |  | 14
20 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.6 | 1.34 | 1.02 | 17.12 | 2520000000 | 5.2069 | 14.2 |  | 15
21 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.5 | 1.31 | 1.01 | 17.56 | 2670000000 | 5.1139 | 14.2 |  | 16
23 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.4 | 1.32 | 1.01 | 17.45 | 2620000000 | 5.0209 | 14.2 |  | 17
24 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.5 | 1.34 | 1.02 | 17.25 | 2330000000 | 5.1139 | 14.2 |  | 18
25 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.4 | 1.33 | 1.01 | 17.56 | 2590000000 | 5.0209 | 14.2 |  | 19
26 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.5 | 1.36 | 1.03 | 17.27 | 2490000000 | 5.1139 | 14.2 |  | 20
28 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.5 | 1.32 | 1.01 | 17.11 | 2430000000 | 5.1139 | 14.2 |  | 21
29 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.4 | 1.33 | 1.02 | 17.48 | 2460000000 | 5.0209 | 14.2 |  | 22
30 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.5 | 1.33 | 1.01 | 17.34 | 2450000000 | 5.1139 | 14.2 |  | 23
31 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.5 | 1.31 | 1 | 17.66 | 2270000000 | 5.1139 | 14.2 |  | 24
32 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.4 | 1.31 | 1.01 | 17.52 | 2430000000 | 5.0209 | 14.2 |  | 25
35 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.4 | 1.33 | 1.01 | 17.57 | 4330000000 | 5.0209 | 14.2 | Re-tested | 26
36 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.4 | 1.3 | 0.99 | 17.72 | 2590000000 | 5.0209 | 14.2 |  | 27
37 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.4 | 1.32 | 1.01 | 17.6 | 2810000000 | 5.0209 | 14.2 |  | 28
38 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.4 | 1.35 | 1.02 | 17.62 | 2540000000 | 5.0209 | 14.2 |  | 29
39 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.6 | 1.31 | 1 | 17.55 | 2610000000 | 5.2069 | 14.2 |  | 30
40 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.5 | 1.34 | 1.02 | 17.1 | 2520000000 | 5.1139 | 14.2 |  | 31
41 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.4 | 1.34 | 1.02 | 17.26 | 2380000000 | 5.0209 | 14.2 |  | 32
42 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.5 | 1.35 | 1.03 | 17.3 | 2550000000 | 5.1139 | 14.2 |  | 33
43 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.4 | 1.29 | 0.99 | 17.75 | 4460000000 | 5.0209 | 14.2 | Re-tested | 34
44 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.4 | 1.31 | 1 | 17.8 | 2540000000 | 5.0209 | 14.8 |  | 35
45 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.5 | 1.32 | 1.01 | 17.6 | 2590000000 | 5.1139 | 14.8 |  | 36
46 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.4 | 1.36 | 1.03 | 17.14 | 3010000000 | 5.0209 | 14.8 |  | 37
47 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.5 | 1.34 | 1.02 | 17.39 | 2540000000 | 5.1139 | 14.8 |  | 38
49 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.4 | 1.34 | 1.02 | 17.46 | 2900000000 | 5.0209 | 14.8 |  | 39
50 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.5 | 1.35 | 1.03 | 17.39 | 2500000000 | 5.1139 | 14.8 |  | 40
51 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.6 | 1.32 | 1.01 | 17.22 | 2760000000 | 5.2069 | 14.8 |  | 41
52 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.4 | 1.29 | 0.99 | 17.69 | 5080000000 | 5.0209 | 14.8 | Re-tested | 42
53 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.6 | 1.35 | 1.02 | 17.46 | 2880000000 | 5.2069 | 14.8 |  | 43
57 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.4 | 1.34 | 1.02 | 17.47 | 2400000000 | 5.0209 | 14.8 |  | 44
59 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.5 | 1.35 | 1.03 | 17.26 | 2520000000 | 5.1139 | 14.8 |  | 45
61 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.6 | 1.36 | 1.03 | 17.29 | 2390000000 | 5.2069 | 14.8 |  | 46
62 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.4 | 1.33 | 1.01 | 17.77 | 2660000000 | 5.0209 | 14.8 |  | 47
63 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.4 | 1.33 | 1.01 | 17.58 | 2400000000 | 5.0209 | 14.8 |  | 48
64 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.4 | 1.32 | 1.01 | 17.47 | 2250000000 | 5.0209 | 14.8 |  | 49
65 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.4 | 1.3 | 1.01 | 17.33 | 2780000000 | 5.0209 | 14.8 |  | 50
66 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.4 | 1.32 | 1.01 | 17.18 | 2630000000 | 5.0209 | 14.8 |  | 51
68 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.4 | 1.32 | 1.01 | 17.35 | 2340000000 | 5.0209 | 14.8 |  | 52
70 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.4 | 1.3 | 1 | 17.47 | 2360000000 | 5.049 | 14.8 |  | 53
71 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.5 | 1.37 | 1.03 | 17.4 | 2540000000 | 5.1425 | 14.8 |  | 54
73 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.4 | 1.31 | 1 | 17.67 | 2170000000 | 5.049 | 14.8 |  | 55
74 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.4 | 1.31 | 1.01 | 17.56 | 2750000000 | 5.049 | 14.8 |  | 56
75 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.6 | 1.31 | 1.01 | 17.38 | 2430000000 | 5.236 | 14.8 |  | 57
76 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.5 | 1.32 | 1.01 | 17.21 | 2950000000 | 5.1425 | 14.8 |  | 58
77 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.5 | 1.32 | 1.02 | 16.95 | 2890000000 | 5.1425 | 14.8 |  | 59
78 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.5 | 1.32 | 1.01 | 17.07 | 2560000000 | 5.1425 | 14.8 |  | 60
79 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.5 | 1.32 | 1.01 | 17.2 | 2550000000 | 5.1425 | 14.8 |  | 61
80 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.3 | 1.31 | 1 | 17.73 | 2410000000 | 4.9555 | 14.8 |  | 62
81 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.3 | 1.31 | 1 | 17.6 | 2850000000 | 4.9555 | 14.8 |  | 63
82 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.4 | 1.3 | 0.99 | 17.54 | 2770000000 | 5.049 | 14.8 |  | 64
83 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.6 | 1.35 | 1.02 | 17.49 | 2490000000 | 5.236 | 14.8 |  | 65
84 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.4 | 1.35 | 1.03 | 17.31 | 2610000000 | 5.049 | 14.8 |  | 66
85 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.3 | 1.33 | 1.01 | 17.55 | 2470000000 | 4.9555 | 14.8 |  | 67
86 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.5 | 1.35 | 1.03 | 17.2 | 2850000000 | 5.1425 | 14.8 |  | 68
87 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.4 | 1.33 | 1.01 | 17.25 | 2560000000 | 5.049 | 14.8 |  | 69
90 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.3 | 1.33 | 1.02 | 17.24 | 2560000000 | 4.9555 | 14.8 |  | 70
91 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.5 | 1.33 | 1.03 | 17.06 | 2370000000 | 5.1425 | 14.8 |  | 71
92 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.4 | 1.33 | 1.02 | 17.27 | 2240000000 | 5.049 | 14.8 |  | 72
93 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.4 | 1.33 | 1.01 | 17.63 | 2190000000 | 5.049 | 14.8 |  | 73
94 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.5 | 1.3 | 0.99 | 17.84 | 1990000000 | 5.1425 | 14.8 |  | 74
95 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.5 | 1.36 | 1.03 | 17.35 | 2540000000 | 5.1425 | 14.8 |  | 75
96 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.4 | 1.33 | 1.02 | 17.45 | 3360000000 | 5.049 | 14.8 |  | 76
97 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.6 | 1.37 | 1.04 | 17.22 | 2820000000 | 5.236 | 14.8 |  | 77
98 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.4 | 1.33 | 1.01 | 17.49 | 2850000000 | 5.049 | 14.8 |  | 78
99 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.4 | 1.32 | 1.01 | 17.39 | 2680000000 | 5.049 | 14.8 |  | 79
100 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.4 | 1.31 | 1 | 17.65 | 2770000000 | 5.049 | 14.8 |  | 80
101 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.3 | 1.31 | 1.01 | 17.54 | 2540000000 | 4.9555 | 14.8 |  | 81
102 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.4 | 1.31 | 1.01 | 17.34 | 4150000000 | 5.049 | 14.8 |  | 82
104 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.4 | 1.36 | 1.03 | 17.13 | 2750000000 | 5.049 | 14.8 |  | 83
105 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.4 | 1.32 | 1.01 | 17.19 | 2460000000 | 5.049 | 14.8 |  | 84
106 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.4 | 1.32 | 1.02 | 17.06 | 2640000000 | 5.049 | 14.8 |  | 85
107 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.4 | 1.34 | 1.03 | 17.14 | 2730000000 | 5.049 | 14.8 |  | 86
108 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.4 | 1.34 | 1.02 | 17.32 | 2840000000 | 5.049 | 14.8 |  | 87
109 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.4 | 1.34 | 1.02 | 17.27 | 2490000000 | 5.049 | 14.8 |  | 88
110 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.4 | 1.37 | 1.04 | 17.17 | 2230000000 | 5.049 | 14.8 |  | 89
111 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.5 | 1.31 | 1.02 | 17.08 | 2460000000 | 5.1425 | 14.8 |  | 90
112 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.5 | 1.35 | 1.04 | 16.95 | 2440000000 | 5.1425 | 14.8 |  | 91
113 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.5 | 1.32 | 1.02 | 17.24 | 2430000000 | 5.1425 | 14.8 |  | 92
115 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.3 | 1.33 | 1.02 | 17.39 | 2510000000 | 4.9555 | 14.8 |  | 93
116 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.5 | 1.31 | 1 | 17.61 | 2320000000 | 5.1425 | 15.3 |  | 94
117 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.3 | 1.31 | 1 | 17.52 | 2650000000 | 4.9555 | 15.3 |  | 95
118 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.4 | 1.31 | 1.01 | 17.17 | 2660000000 | 5.049 | 15.3 |  | 96
119 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.5 | 1.35 | 1.03 | 17.04 | 2750000000 | 5.1425 | 15.3 |  | 97
120 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.4 | 1.34 | 1.03 | 17.22 | 2700000000 | 5.049 | 15.3 |  | 98
121 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.3 | 1.34 | 1.03 | 17.32 | 2380000000 | 4.9555 | 15.3 |  | 99
122 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.4 | 1.3 | 1.01 | 17.06 | 2650000000 | 5.049 | 15.3 |  | 100
123 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.3 | 1.34 | 1.02 | 17.31 | 2470000000 | 4.9555 | 15.3 |  | 101
124 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.4 | 1.34 | 1.02 | 17.22 | 2400000000 | 5.049 | 15.3 |  | 102
125 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.4 | 1.34 | 1.01 | 17.42 | 2520000000 | 5.049 | 15.3 |  | 103
126 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.5 | 1.37 | 1.04 | 17.11 | 2510000000 | 5.1425 | 15.3 |  | 104
127 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.3 | 1.35 | 1.04 | 16.91 | 2600000000 | 4.9555 | 15.3 |  | 105
128 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.5 | 1.38 | 1.05 | 16.81 | 2380000000 | 5.1425 | 15.3 |  | 106
131 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.5 | 1.33 | 1.02 | 17.37 | 2390000000 | 5.1425 | 15.3 |  | 107
132 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.6 | 1.35 | 1.02 | 17.44 | 2710000000 | 5.236 | 15.3 |  | 108
133 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.5 | 1.35 | 1.02 | 17.29 | 2680000000 | 5.1425 | 15.3 |  | 109
135 | 1087 | XIRE0313MS6 |  | x | 21.5 | 5.6 | 1.31 | 1.01 | 17.08 | 2680000000 | 5.236 | 15.3 |  | 110
136 | 1087 | XIRE0313MS6 |  | x | 21.6 | 5.4 | 1.34 | 1.02 | 17.19 | 3380000000 | 5.063 | 15.4 |  | 111
137 | 1087 | XIRE0313MS6 |  | x | 21.6 | 5.3 | 1.31 | 1.02 | 16.96 | 2540000000 | 4.9693 | 15.4 |  | 112
138 | 1087 | XIRE0313MS6 |  | x | 21.6 | 5.5 | 1.36 | 1.04 | 17.11 | 2190000000 | 5.1568 | 15.4 |  | 113
139 | 1087 | XIRE0313MS6 |  | x | 21.6 | 5.3 | 1.31 | 1.01 | 17.37 | 2520000000 | 4.9693 | 15.4 |  | 114
140 | 1087 | XIRE0313MS6 |  | x | 21.6 | 5.4 | 1.36 | 1.04 | 17.06 | 2310000000 | 5.063 | 15.4 |  | 115
141 | 1087 | XIRE0313MS6 |  | x | 21.6 | 5.5 | 1.36 | 1.05 | 16.84 | 2730000000 | 5.1568 | 15.4 |  | 116
142 | 1087 | XIRE0313MS6 |  | x | 21.6 | 5.5 | 1.33 | 1.03 | 17 | 2530000000 | 5.1568 | 15.4 |  | 117
143 | 1087 | XIRE0313MS6 |  | x | 21.6 | 5.4 | 1.33 | 1.02 | 17.35 | 3070000000 | 5.063 | 15.4 |  | 118
144 | 1087 | XIRE0313MS6 |  | x | 21.6 | 5.4 | 1.35 | 1.03 | 17.26 | 2680000000 | 5.063 | 15.4 |  | 119
145 | 1087 | XIRE0313MS6 |  | x | 21.6 | 5.4 | 1.39 | 1.06 | 16.89 | 2480000000 | 5.063 | 15.4 |  | 120
147 | 1087 | XIRE0313MS6 |  | x | 21.6 | 5.3 | 1.33 | 1.03 | 17.07 | 2750000000 | 4.9693 | 15.4 |  | 121
148 | 1087 | XIRE0313MS6 |  | x | 21.6 | 5.3 | 1.33 | 1.03 | 17.09 | 2900000000 | 4.9693 | 15.4 |  | 122
149 | 1087 | XIRE0313MS6 |  | x | 21.6 | 5.5 | 1.33 | 1.02 | 16.91 | 3000000000 | 5.1568 | 15.4 |  | 123
150 | 1087 | XIRE0313MS6 |  | x | 21.6 | 5.4 | 1.33 | 1.03 | 16.92 | 2540000000 | 5.063 | 15.4 |  | 124
151 | 1087 | XIRE0313MS6 |  | x | 21.6 | 5.4 | 1.33 | 1.04 | 16.8 | 2580000000 | 5.063 | 15.4 |  | 125
153 | 1087 | XIRE0313MS6 |  | x | 21.6 | 5.4 | 1.35 | 1.04 | 16.74 | 2640000000 | 5.063 | 15.4 |  | 126
154 | 1087 | XIRE0313MS6 |  | x | 21.6 | 5.4 | 1.33 | 1.03 | 17.01 | 3010000000 | 5.063 | 15.4 |  | 127
155 | 1087 | XIRE0313MS6 |  | x | 21.6 | 5.4 | 1.35 | 1.04 | 17.07 | 2850000000 | 5.063 | 15.4 |  | 128
156 | 1087 | XIRE0313MS6 |  | x | 21.6 | 5.3 | 1.35 | 1.04 | 17.09 | 2640000000 | 4.9693 | 15.4 |  | 129
157 | 1087 | XIRE0313MS6 |  | x | 21.6 | 5.5 | 1.37 | 1.05 | 16.99 | 2550000000 | 5.1568 | 15.4 |  | 130
158 | 1087 | XIRE0313MS6 |  | x | 21.6 | 5.4 | 1.34 | 1.04 | 16.8 | 2590000000 | 5.063 | 15.4 |  | 131
159 | 1087 | XIRE0313MS6 |  | x | 21.6 | 5.4 | 1.36 | 1.04 | 17.12 | 2540000000 | 5.063 | 15.4 |  | 132
161 | 1087 | XIRE0313MS6 |  | x | 21.6 | 5.3 | 1.31 | 1.02 | 16.97 | 2670000000 | 4.9693 | 15.4 |  | 133
162 | 1087 | XIRE0313MS6 |  | x | 21.6 | 5.3 | 1.35 | 1.03 | 17.07 | 2650000000 | 4.9693 | 15.4 |  | 134
163 | 1087 | XIRE0313MS6 |  | x | 21.6 | 5.4 | 1.33 | 1.03 | 16.94 | 2860000000 | 5.063 | 15.4 |  | 135
164 | 1087 | XIRE0313MS6 |  | x | 21.6 | 5.4 | 1.33 | 1.03 | 16.87 | 2960000000 | 5.063 | 15.4 |  | 136
165 | 1087 | XIRE0313MS6 |  | x | 21.6 | 5.3 | 1.33 | 1.02 | 17.05 | 2850000000 | 4.9693 | 15.4 |  | 137
166 | 1087 | XIRE0313MS6 |  | x | 21.6 | 5.5 | 1.31 | 1.01 | 16.82 | 6230000000 | 5.1568 | 15.4 |  | 138
167 | 1087 | XIRE0313MS6 |  | x | 21.6 | 5.5 | 1.35 | 1.05 | 16.48 | 2680000000 | 5.1568 | 15.4 |  | 139
168 | 1087 | XIRE0313MS6 |  | x | 21.6 | 5.4 | 1.33 | 1.02 | 17.1 | 2560000000 | 5.063 | 15.4 |  | 140
169 | 1087 | XIRE0313MS6 |  | x | 21.6 | 5.3 | 1.36 | 1.03 | 17.12 | 2570000000 | 4.9693 | 15.4 |  | 141
170 | 1087 | XIRE0313MS6 |  | x | 21.6 | 5.3 | 1.32 | 1.02 | 17.16 | 2750000000 | 4.9693 | 15.4 |  | 142
171 | 1087 | XIRE0313MS6 |  | x | 21.6 | 5.3 | 1.3 | 1.01 | 17.13 | 2420000000 | 4.9693 | 15.4 |  | 143
172 | 1087 | XIRE0313MS6 |  | x | 21.7 | 5.5 | 1.33 | 1.02 | 17.08 | 3080000000 | 5.1711 | 16.5 |  | 144
173 | 1087 | XIRE0313MS6 |  | x | 21.7 | 5.4 | 1.33 | 1.02 | 16.87 | 2730000000 | 5.0771 | 16.5 |  | 145
176 | 1087 | XIRE0313MS6 |  | x | 21.7 | 5.4 | 1.34 | 1.02 | 17.19 | 2990000000 | 5.0771 | 16.5 |  | 146
177 | 1087 | XIRE0313MS6 |  | x | 21.7 | 5.3 | 1.31 | 1.01 | 17.29 | 2660000000 | 4.9831 | 16.5 |  | 147
178 | 1087 | XIRE0313MS6 |  | x | 21.7 | 5.3 | 1.34 | 1.03 | 17.06 | 2790000000 | 4.9831 | 16.5 |  | 148
179 | 1087 | XIRE0313MS6 |  | x | 21.7 | 5.3 | 1.32 | 1.02 | 17.38 | 2760000000 | 4.9831 | 16.5 |  | 149
180 | 1087 | XIRE0313MS6 |  | x | 21.7 | 5.4 | 1.34 | 1.03 | 16.82 | 2320000000 | 5.0771 | 16.5 |  | 150
182 | 1087 | XIRE0313MS6 |  | x | 21.7 | 5.5 | 1.34 | 1.03 | 16.91 | 2370000000 | 5.1711 | 16.5 |  | 151
183 | 1087 | XIRE0313MS6 |  | x | 21.7 | 5.4 | 1.34 | 1.02 | 17.19 | 2320000000 | 5.0771 | 16.5 |  | 152
184 | 1087 | XIRE0313MS6 |  | x | 21.7 | 5.3 | 1.36 | 1.03 | 17.19 | 2680000000 | 4.9831 | 16.5 |  | 153
185 | 1087 | XIRE0313MS6 |  | x | 21.7 | 5.3 | 1.33 | 1.02 | 17.1 | 2730000000 | 4.9831 | 16.5 |  | 154
187 | 1087 | XIRE0313MS6 |  | x | 21.7 | 5.4 | 1.35 | 1.03 | 17.15 | 2650000000 | 5.0771 | 16.5 |  | 155
188 | 1087 | XIRE0313MS6 |  | x | 21.7 | 5.3 | 1.35 | 1.03 | 17.11 | 2950000000 | 4.9831 | 16.5 |  | 156
189 | 1087 | XIRE0313MS6 |  | x | 21.7 | 5.4 | 1.33 | 1.03 | 17.03 | 2320000000 | 5.0771 | 16.5 |  | 157
190 | 1087 | XIRE0313MS6 |  | x | 21.7 | 5.3 | 1.33 | 1.02 | 17.33 | 2950000000 | 4.9831 | 16.5 |  | 158
191 | 1087 | XIRE0313MS6 |  | x | 21.7 | 5.3 | 1.33 | 1.02 | 17.21 | 3010000000 | 4.9831 | 16.5 |  | 159
192 | 1087 | XIRE0313MS6 |  | x | 21.7 | 5.5 | 1.32 | 1.05 | 16.82 | 2600000000 | 5.1711 | 16.5 |  | 160
193 | 1087 | XIRE0313MS6 |  | x | 21.7 | 5.4 | 1.32 | 1.02 | 16.82 | 2840000000 | 5.0771 | 16.5 |  | 161
196 | 1087 | XIRE0313MS6 |  | x | 21.8 | 5.4 | 1.34 | 1.03 | 17.03 | 2800000000 | 5.0911 | 18.4 |  | 162
197 | 1087 | XIRE0313MS6 |  | x | 21.8 | 5.3 | 1.31 | 1.01 | 17.26 | 3450000000 | 4.9968 | 18.4 |  | 163
198 | 1087 | XIRE0313MS6 |  | x | 21.8 | 5.5 | 1.32 | 1.02 | 16.74 | 2980000000 | 5.1854 | 18.4 |  | 164
199 | 1087 | XIRE0313MS6 |  | x | 21.8 | 5.5 | 1.36 | 1.05 | 16.85 | 2820000000 | 5.1854 | 18.4 |  | 165
200 | 1087 | XIRE0313MS6 |  | x | 21.8 | 5.4 | 1.32 | 1.03 | 16.9 | 2490000000 | 5.0911 | 18.4 |  | 166
201 | 1087 | XIRE0313MS6 |  | x | 21.8 | 5.5 | 1.36 | 1.04 | 16.8 | 2480000000 | 5.1854 | 18.4 |  | 167
202 | 1087 | XIRE0313MS6 |  | x | 21.8 | 5.3 | 1.31 | 1.01 | 17.36 | 3140000000 | 4.9968 | 18.4 |  | 168
203 | 1087 | XIRE0313MS6 |  | x | 21.8 | 5.5 | 1.35 | 1.03 | 17.09 | 2530000000 | 5.1854 | 18.4 |  | 169
204 | 1087 | XIRE0313MS6 |  | x | 21.8 | 5.3 | 1.35 | 1.04 | 17.11 | 2540000000 | 4.9968 | 18.4 |  | 170
205 | 1087 | XIRE0313MS6 |  | x | 21.8 | 5.3 | 1.35 | 1.03 | 17.04 | 2610000000 | 4.9968 | 18.4 |  | 171
206 | 1087 | XIRE0313MS6 |  | x | 21.8 | 5.5 | 1.35 | 1.04 | 16.81 | 2400000000 | 5.1854 | 18.4 |  | 172
207 | 1087 | XIRE0313MS6 |  | x | 21.8 | 5.4 | 1.35 | 1.04 | 17.09 | 2100000000 | 5.0911 | 18.4 |  | 173
208 | 1087 | XIRE0313MS6 |  | x | 21.8 | 5.5 | 1.35 | 1.04 | 16.65 | 3210000000 | 5.1854 | 18.4 |  | 174
209 | 1087 | XIRE0313MS6 |  | x | 21.8 | 5.3 | 1.3 | 1.01 | 16.83 | 2590000000 | 4.9968 | 18.4 |  | 175
210 | 1087 | XIRE0313MS6 |  | x | 21.8 | 5.3 | 1.34 | 1.03 | 16.76 | 2680000000 | 4.9968 | 18.4 |  | 176
211 | 1087 | XIRE0313MS6 |  | x | 21.8 | 5.4 | 1.3 | 1.02 | 16.62 | 2610000000 | 5.0911 | 18.4 |  | 177
213 | 1087 | XIRE0313MS6 |  | x | 21.8 | 5.3 | 1.34 | 1.03 | 17.02 | 3110000000 | 4.9968 | 18.4 |  | 178
214 | 1087 | XIRE0313MS6 |  | x | 21.8 | 5.4 | 1.33 | 1.03 | 17.04 | 3420000000 | 5.0911 | 18.4 |  | 179
215 | 1087 | XIRE0313MS6 |  | x | 21.8 | 5.3 | 1.3 | 1.01 | 16.88 | 3160000000 | 4.9968 | 18.4 |  | 180
219 | 1087 | XIRE0313MS6 |  | x | 21.8 | 5.4 | 1.32 | 1.04 | 16.36 | 3010000000 | 5.0911 | 18.4 |  | 181
221 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.4 | 1.36 | 1.05 | 16.74 | 7000000000 | 5.1052 | 19.5 | Re-tested | 182
223 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.3 | 1.33 | 1.03 | 16.51 | 3240000000 | 5.0106 | 19.5 |  | 183
224 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.3 | 1.33 | 1.03 | 16.78 | 3180000000 | 5.0106 | 19.5 |  | 184
225 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.4 | 1.36 | 1.04 | 16.83 | 3090000000 | 5.1052 | 19.5 |  | 185
226 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.3 | 1.32 | 1.02 | 16.88 | 2840000000 | 5.0106 | 19.5 |  | 186
227 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.5 | 1.37 | 1.06 | 16.6 | 2980000000 | 5.1997 | 19.5 |  | 187
228 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.4 | 1.36 | 1.05 | 16.78 | 3130000000 | 5.1052 | 19.5 |  | 188
229 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.3 | 1.32 | 1.02 | 16.84 | 2810000000 | 5.0106 | 19.5 |  | 189
230 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.4 | 1.35 | 1.04 | 16.64 | 3760000000 | 5.1052 | 19.5 |  | 190
232 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.4 | 1.3 | 1.02 | 17.04 | 2860000000 | 5.1052 | 19.5 |  | 191
233 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.5 | 1.32 | 1.03 | 16.85 | 3110000000 | 5.1997 | 19.5 |  | 192
236 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.4 | 1.35 | 1.05 | 16.3 | 3220000000 | 5.1052 | 19.5 |  | 193
237 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.3 | 1.32 | 1.03 | 16.69 | 3090000000 | 5.0106 | 19.5 |  | 194
238 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.4 | 1.35 | 1.05 | 16.65 | 3010000000 | 5.1052 | 19.5 |  | 195
239 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.3 | 1.35 | 1.04 | 16.75 | 3350000000 | 5.0106 | 19.5 |  | 196
241 | 1087 | XIRE0313MS6 |  | x | 22 | 5.4 | 1.35 | 1.04 | 16.75 | 2540000000 | 5.1192 | 20.3 |  | 197
242 | 1087 | XIRE0313MS6 |  | x | 22 | 5.5 | 1.35 | 1.04 | 16.59 | 3420000000 | 5.214 | 20.3 |  | 198
243 | 1087 | XIRE0313MS6 |  | x | 22 | 5.3 | 1.31 | 1.02 | 17.07 | 2910000000 | 5.0244 | 20.3 |  | 199
244 | 1087 | XIRE0313MS6 |  | x | 22 | 5.4 | 1.34 | 1.04 | 16.9 | 2810000000 | 5.1192 | 20.3 |  | 200
245 | 1087 | XIRE0313MS6 |  | x | 22 | 5.5 | 1.36 | 1.05 | 16.46 | 3410000000 | 5.214 | 20.3 |  | 201
246 | 1087 | XIRE0313MS6 |  | x | 22 | 5.5 | 1.35 | 1.04 | 16.67 | 2820000000 | 5.214 | 20.3 |  | 202
247 | 1087 | XIRE0313MS6 |  | x | 22 | 5.5 | 1.35 | 1.05 | 16.26 | 2720000000 | 5.214 | 20.3 |  | 203
248 | 1087 | XIRE0313MS6 |  | x | 22 | 5.5 | 1.35 | 1.05 | 16.43 | 2520000000 | 5.214 | 20.3 |  | 204
249 | 1087 | XIRE0313MS6 |  | x | 22 | 5.3 | 1.31 | 1.01 | 17.04 | 2630000000 | 5.0244 | 20.3 |  | 205
250 | 1087 | XIRE0313MS6 |  | x | 22 | 5.3 | 1.3 | 1.01 | 16.97 | 2850000000 | 5.0244 | 20.3 |  | 206
252 | 1087 | XIRE0313MS6 |  | x | 22 | 5.3 | 1.33 | 1.02 | 17.11 | 2750000000 | 5.0244 | 20.3 |  | 207
253 | 1087 | XIRE0313MS6 |  | x | 22 | 5.4 | 1.36 | 1.05 | 16.62 | 2930000000 | 5.1192 | 20.3 |  | 208
255 | 1087 | XIRE0313MS6 |  | x | 22 | 5.4 | 1.36 | 1.04 | 16.79 | 2490000000 | 5.1192 | 20.3 |  | 209
256 | 1087 | XIRE0313MS6 |  | x | 22 | 5.4 | 1.33 | 1.03 | 17.13 | 6470000000 | 5.1192 | 20.3 | Re-tested | 210
257 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.4 | 1.33 | 1.03 | 17 | 2540000000 | 5.1052 | 20.6 |  | 211
261 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.5 | 1.35 | 1.05 | 16.57 | 2820000000 | 5.1997 | 20.6 |  | 212
262 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.5 | 1.37 | 1.05 | 16.48 | 2550000000 | 5.1997 | 20.6 |  | 213
263 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.4 | 1.34 | 1.03 | 16.73 | 2740000000 | 5.1052 | 20.6 |  | 214
265 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.4 | 1.33 | 1.03 | 16.73 | 2630000000 | 5.1052 | 20.6 |  | 215
266 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.4 | 1.34 | 1.04 | 17 | 2320000000 | 5.1052 | 20.6 |  | 216
267 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.5 | 1.37 | 1.05 | 16.62 | 2550000000 | 5.1997 | 20.6 |  | 217
269 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.4 | 1.38 | 1.06 | 16.63 | 2850000000 | 5.1052 | 20.6 |  | 218
270 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.3 | 1.34 | 1.04 | 16.68 | 2590000000 | 5.0106 | 20.6 |  | 219
271 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.3 | 1.33 | 1.03 | 16.73 | 2310000000 | 5.0106 | 20.6 |  | 220
272 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.4 | 1.33 | 1.03 | 16.73 | 2390000000 | 5.1052 | 20.6 |  | 221
274 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.3 | 1.37 | 1.06 | 16.45 | 2500000000 | 5.0106 | 20.6 |  | 222
275 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.4 | 1.37 | 1.05 | 16.56 | 2640000000 | 5.1052 | 20.6 |  | 223
276 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.3 | 1.38 | 1.06 | 16.47 | 2380000000 | 5.0106 | 20.6 |  | 224
277 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.3 | 1.32 | 1.02 | 16.9 | 2670000000 | 5.0106 | 20.6 |  | 225
278 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.3 | 1.32 | 1.03 | 16.7 | 2220000000 | 5.0106 | 20.6 |  | 226
279 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.4 | 1.37 | 1.06 | 16.36 | 2650000000 | 5.1052 | 20.6 |  | 227
280 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.4 | 1.37 | 1.05 | 16.7 | 2400000000 | 5.1052 | 20.6 |  | 228
281 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.4 | 1.34 | 1.04 | 16.67 | 2770000000 | 5.1052 | 20.6 |  | 229
285 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.5 | 1.39 | 1.07 | 16.33 | 2260000000 | 5.1997 | 20.6 |  | 230
287 | 1087 | XIRE0313MS6 |  | x | 22 | 5.3 | 1.35 | 1.04 | 17.01 | 3000000000 | 5.0244 | 20.9 |  | 231
290 | 1087 | XIRE0313MS6 |  | x | 22 | 5.3 | 1.34 | 1.04 | 16.64 | 2930000000 | 5.0244 | 20.9 |  | 232
291 | 1087 | XIRE0313MS6 |  | x | 22 | 5.4 | 1.34 | 1.04 | 16.66 | 5640000000 | 5.1192 | 20.9 |  | 233
293 | 1087 | XIRE0313MS6 |  | x | 22 | 5.3 | 1.32 | 1.03 | 16.78 | 2490000000 | 5.0244 | 20.9 |  | 234
294 | 1087 | XIRE0313MS6 |  | x | 22 | 5.3 | 1.32 | 1.02 | 16.87 | 2820000000 | 5.0244 | 20.9 |  | 235
295 | 1087 | XIRE0313MS6 |  | x | 22 | 5.4 | 1.34 | 1.04 | 16.66 | 2620000000 | 5.1192 | 20.9 |  | 236
296 | 1087 | XIRE0313MS6 |  | x | 22 | 5.4 | 1.36 | 1.06 | 16.22 | 2510000000 | 5.1192 | 20.9 |  | 237
297 | 1087 | XIRE0313MS6 |  | x | 22 | 5.4 | 1.33 | 1.03 | 16.58 | 2850000000 | 5.1192 | 20.9 |  | 238
299 | 1087 | XIRE0313MS6 |  | x | 22 | 5.4 | 1.35 | 1.04 | 16.68 | 2640000000 | 5.1192 | 20.9 |  | 239
301 | 1087 | XIRE0313MS6 |  | x | 22 | 5.5 | 1.37 | 1.05 | 16.36 | 2250000000 | 5.214 | 20.9 |  | 240
302 | 1087 | XIRE0313MS6 |  | x | 22 | 5.4 | 1.36 | 1.05 | 16.64 | 2350000000 | 5.1192 | 20.9 |  | 241
304 | 1087 | XIRE0313MS6 |  | x | 22 | 5.5 | 1.38 | 1.06 | 16.69 | 2650000000 | 5.214 | 20.9 |  | 242
305 | 1087 | XIRE0313MS6 |  | x | 22 | 5.4 | 1.34 | 1.04 | 16.63 | 3190000000 | 5.1192 | 20.9 |  | 243
306 | 1087 | XIRE0313MS6 |  | x | 22 | 5.3 | 1.34 | 1.03 | 17.19 | 2840000000 | 5.0244 | 20.9 |  | 244
307 | 1087 | XIRE0313MS6 |  | x | 22 | 5.4 | 1.32 | 1.04 | 16.5 | 2660000000 | 5.1192 | 20.9 |  | 245
308 | 1087 | XIRE0313MS6 |  | x | 22 | 5.4 | 1.34 | 1.05 | 16.4 | 3110000000 | 5.1192 | 20.9 |  | 246
309 | 1087 | XIRE0313MS6 |  | x | 22 | 5.3 | 1.32 | 1.03 | 16.66 | 2770000000 | 5.0244 | 20.9 |  | 247
310 | 1087 | XIRE0313MS6 |  | x | 22 | 5.5 | 1.36 | 1.05 | 16.61 | 3080000000 | 5.214 | 20.9 |  | 248
311 | 1087 | XIRE0313MS6 |  | x | 22 | 5.5 | 1.36 | 1.05 | 16.4 | 4560000000 | 5.214 | 20.9 |  | 249
313 | 1087 | XIRE0313MS6 |  | x | 22 | 5.4 | 1.37 | 1.06 | 16.57 | 2580000000 | 5.1192 | 20.9 |  | 250
314 | 1087 | XIRE0313MS6 |  | x | 22 | 5.3 | 1.31 | 1.01 | 16.88 | 2480000000 | 5.0244 | 20.9 |  | 251
316 | 1087 | XIRE0313MS6 |  | x | 22 | 5.4 | 1.34 | 1.04 | 16.78 | 3380000000 | 5.1192 | 20.9 |  | 252
317 | 1087 | XIRE0313MS6 |  | x | 22 | 5.3 | 1.34 | 1.03 | 17.03 | 3180000000 | 5.0244 | 21 |  | 253
318 | 1087 | XIRE0313MS6 |  | x | 22 | 5.3 | 1.33 | 1.03 | 16.85 | 2750000000 | 5.0244 | 21 |  | 254
319 | 1087 | XIRE0313MS6 |  | x | 22 | 5.3 | 1.33 | 1.03 | 16.65 | 2300000000 | 5.0244 | 21 |  | 255
320 | 1087 | XIRE0313MS6 |  | x | 22 | 5.5 | 1.39 | 1.07 | 16.34 | 2780000000 | 5.214 | 21 |  | 256
321 | 1087 | XIRE0313MS6 |  | x | 22 | 5.3 | 1.37 | 1.06 | 16.42 | 2150000000 | 5.0244 | 21 |  | 257
322 | 1087 | XIRE0313MS6 |  | x | 22 | 5.3 | 1.35 | 1.05 | 16.47 | 2810000000 | 5.0244 | 21 |  | 258
323 | 1087 | XIRE0313MS6 |  | x | 22 | 5.4 | 1.38 | 1.06 | 16.81 | 2990000000 | 5.1192 | 21 |  | 259
324 | 1087 | XIRE0313MS6 |  | x | 22 | 5.3 | 1.36 | 1.05 | 16.55 | 2440000000 | 5.0244 | 21 |  | 260
325 | 1087 | XIRE0313MS6 |  | x | 22 | 5.4 | 1.33 | 1.02 | 16.59 | 7300000000 | 5.1192 | 21 |  | 261
326 | 1087 | XIRE0313MS6 |  | x | 22 | 5.4 | 1.36 | 1.05 | 16.84 | 2700000000 | 5.1192 | 21 |  | 262
327 | 1087 | XIRE0313MS6 |  | x | 22 | 5.4 | 1.32 | 1.03 | 16.43 | 2480000000 | 5.1192 | 21 |  | 263
328 | 1087 | XIRE0313MS6 |  | x | 22 | 5.3 | 1.35 | 1.04 | 16.48 | 3050000000 | 5.0244 | 21 |  | 264
330 | 1087 | XIRE0313MS6 |  | x | 22 | 5.5 | 1.34 | 1.04 | 16.73 | 2730000000 | 5.214 | 21 |  | 265
331 | 1087 | XIRE0313MS6 |  | x | 22 | 5.4 | 1.34 | 1.04 | 16.61 | 2850000000 | 5.1192 | 21 |  | 266
332 | 1087 | XIRE0313MS6 |  | x | 22 | 5.4 | 1.34 | 1.04 | 16.45 | 2370000000 | 5.1192 | 21 |  | 267
333 | 1087 | XIRE0313MS6 |  | x | 22 | 5.4 | 1.33 | 1.04 | 16.42 | 2620000000 | 5.1192 | 21 |  | 268
334 | 1087 | XIRE0313MS6 |  | x | 22 | 5.4 | 1.37 | 1.05 | 16.87 | 1850000000 | 5.1192 | 21 |  | 269
335 | 1087 | XIRE0313MS6 |  | x | 22 | 5.4 | 1.33 | 1.03 | 16.56 | 2850000000 | 5.1192 | 21 |  | 270
336 | 1087 | XIRE0313MS6 |  | x | 22 | 5.4 | 1.33 | 1.04 | 16.63 | 2720000000 | 5.1192 | 21 |  | 271
337 | 1087 | XIRE0313MS6 |  | x | 22 | 5.4 | 1.32 | 1.03 | 16.56 | 2760000000 | 5.1192 | 21 |  | 272
338 | 1087 | XIRE0313MS6 |  | x | 22 | 5.4 | 1.36 | 1.05 | 16.56 | 2820000000 | 5.1192 | 21 |  | 273
339 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.3 | 1.34 | 1.04 | 16.49 | 2750000000 | 5.0382 | 21.5 |  | 274
340 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.4 | 1.36 | 1.05 | 16.59 | 8810000000 | 5.1332 | 21.5 | Re-tested | 275
341 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.5 | 1.35 | 1.06 | 16.42 | 3070000000 | 5.2283 | 21.5 |  | 276
342 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.5 | 1.32 | 1.02 | 16.83 | 2770000000 | 5.2283 | 21.5 |  | 277
343 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.5 | 1.33 | 1.04 | 16.29 | 2850000000 | 5.2283 | 21.5 |  | 278
344 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.4 | 1.36 | 1.05 | 16.77 | 2570000000 | 5.1332 | 21.5 |  | 279
345 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.4 | 1.36 | 1.05 | 16.62 | 3130000000 | 5.1332 | 21.5 |  | 280
347 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.5 | 1.33 | 1.03 | 16.52 | 2540000000 | 5.2283 | 21.5 |  | 281
349 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.4 | 1.33 | 1.03 | 17.28 | 2930000000 | 5.1332 | 21.5 |  | 282
350 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.3 | 1.34 | 1.04 | 16.94 | 3110000000 | 5.0382 | 21.5 |  | 283
351 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.3 | 1.34 | 1.03 | 16.97 | 2710000000 | 5.0382 | 21.5 |  | 284
352 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.4 | 1.34 | 1.04 | 16.63 | 2520000000 | 5.1332 | 21.5 |  | 285
353 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.5 | 1.33 | 1.04 | 16.55 | 2790000000 | 5.2283 | 21.5 |  | 286
354 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.4 | 1.33 | 1.04 | 16.42 | 3050000000 | 5.1332 | 21.5 |  | 287
355 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.3 | 1.32 | 1.03 | 16.56 | 3320000000 | 5.0382 | 21.5 |  | 288
356 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.3 | 1.33 | 1.03 | 17.04 | 2480000000 | 5.0382 | 21.5 |  | 289
357 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.3 | 1.32 | 1.02 | 16.95 | 3160000000 | 5.0382 | 21.5 |  | 290
358 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.3 | 1.33 | 1.03 | 16.87 | 2970000000 | 5.0382 | 21.5 |  | 291
360 | 1087 | XIRE0313MS6 |  | x | 22 | 5.3 | 1.33 | 1.03 | 16.8 | 3350000000 | 5.0244 | 21.4 |  | 292
361 | 1087 | XIRE0313MS6 |  | x | 22 | 5.4 | 1.34 | 1.04 | 17.02 | 2780000000 | 5.1192 | 21.4 |  | 293
364 | 1087 | XIRE0313MS6 |  | x | 22 | 5.5 | 1.35 | 1.05 | 16.38 | 2850000000 | 5.214 | 21.4 |  | 294
365 | 1087 | XIRE0313MS6 |  | x | 22 | 5.3 | 1.3 | 1.01 | 17.02 | 2900000000 | 5.0244 | 21.4 |  | 295
366 | 1087 | XIRE0313MS6 |  | x | 22 | 5.5 | 1.35 | 1.05 | 16.43 | 3200000000 | 5.214 | 21.4 |  | 296
367 | 1087 | XIRE0313MS6 |  | x | 22 | 5.4 | 1.31 | 1.01 | 17.11 | 3620000000 | 5.1192 | 21.4 |  | 297
368 | 1087 | XIRE0313MS6 |  | x | 22 | 5.3 | 1.33 | 1.03 | 16.7 | 2950000000 | 5.0244 | 21.4 |  | 298
369 | 1087 | XIRE0313MS6 |  | x | 22 | 5.3 | 1.36 | 1.05 | 16.72 | 2590000000 | 5.0244 | 21.4 |  | 299
370 | 1087 | XIRE0313MS6 |  | x | 22 | 5.5 | 1.34 | 1.04 | 16.82 | 2950000000 | 5.214 | 21.4 |  | 300
371 | 1087 | XIRE0313MS6 |  | x | 22 | 5.4 | 1.33 | 1.03 | 16.68 | 2340000000 | 5.1192 | 21.4 |  | 301
372 | 1087 | XIRE0313MS6 |  | x | 22 | 5.3 | 1.35 | 1.04 | 16.8 | 2930000000 | 5.0244 | 21.4 |  | 302
375 | 1087 | XIRE0313MS6 |  | x | 22 | 5.4 | 1.34 | 1.04 | 16.5 | 2950000000 | 5.1192 | 21.4 |  | 303
376 | 1087 | XIRE0313MS6 |  | x | 22 | 5.3 | 1.31 | 1.02 | 16.87 | 2820000000 | 5.0244 | 21.4 |  | 304
377 | 1087 | XIRE0313MS6 |  | x | 22 | 5.5 | 1.34 | 1.04 | 16.92 | 2630000000 | 5.214 | 21.4 | Re-tested | 305
378 | 1087 | XIRE0313MS6 |  | x | 22 | 5.5 | 1.34 | 1.04 | 16.62 | 2850000000 | 5.214 | 21.4 |  | 306
379 | 1087 | XIRE0313MS6 |  | x | 22 | 5.5 | 1.33 | 1.03 | 16.7 | 3140000000 | 5.214 | 21.4 |  | 307
380 | 1087 | XIRE0313MS6 |  | x | 22 | 5.4 | 1.34 | 1.04 | 16.77 | 3250000000 | 5.1192 | 21.4 |  | 308
381 | 1087 | XIRE0313MS6 |  | x | 22 | 5.3 | 1.33 | 1.02 | 16.92 | 2980000000 | 5.0244 | 21.4 |  | 309
382 | 1087 | XIRE0313MS6 |  | x | 22 | 5.3 | 1.33 | 1.03 | 16.99 | 2950000000 | 5.0244 | 21.4 |  | 310
384 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.3 | 1.3 | 1.01 | 16.87 | 3030000000 | 5.0106 | 22.5 |  | 311
386 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.5 | 1.39 | 1.07 | 16.47 | 2990000000 | 5.1997 | 22.5 |  | 312
387 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.4 | 1.36 | 1.04 | 16.74 | 2570000000 | 5.1052 | 22.5 |  | 313
388 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.4 | 1.35 | 1.04 | 17.09 | 2900000000 | 5.1052 | 22.5 |  | 314
389 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.3 | 1.31 | 1.01 | 16.9 | 2880000000 | 5.0106 | 22.5 |  | 315
390 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.4 | 1.35 | 1.04 | 16.87 | 2870000000 | 5.1052 | 22.5 |  | 316
391 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.4 | 1.34 | 1.04 | 16.98 | 3210000000 | 5.1052 | 22.5 |  | 317
393 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.4 | 1.35 | 1.04 | 16.92 | 2830000000 | 5.1052 | 22.5 |  | 318
394 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.3 | 1.33 | 1.03 | 17.09 | 5240000000 | 5.0106 | 22.5 | Re-tested | 319
395 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.4 | 1.33 | 1.02 | 16.96 | 2500000000 | 5.1052 | 22.5 |  | 320
396 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.4 | 1.34 | 1.03 | 17.05 | 3190000000 | 5.1052 | 22.5 |  | 321
397 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.4 | 1.33 | 1.03 | 17.15 | 2900000000 | 5.1052 | 22.5 |  | 322
398 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.4 | 1.34 | 1.03 | 17 | 3290000000 | 5.1052 | 22.5 |  | 323
399 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.3 | 1.32 | 1.03 | 16.87 | 2820000000 | 5.0106 | 22.5 |  | 324
400 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.5 | 1.35 | 1.04 | 16.52 | 2640000000 | 5.1997 | 22.5 |  | 325
401 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.4 | 1.34 | 1.04 | 16.85 | 2670000000 | 5.1052 | 22.5 |  | 326
402 | 1087 | XIRE0313MS6 |  | x | 21.9 | 5.4 | 1.35 | 1.04 | 16.89 | 3680000000 | 5.1052 | 22.5 | Re-tested | 327
403 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.3 | 1.33 | 1.03 | 17.03 | 3140000000 | 5.0382 | 21.4 |  | 328
404 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.4 | 1.31 | 1.02 | 16.84 | 2820000000 | 5.1332 | 21.4 |  | 329
406 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.2 | 1.31 | 1.02 | 16.97 | 3350000000 | 4.9431 | 21.4 |  | 330
407 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.4 | 1.35 | 1.04 | 16.78 | 2750000000 | 5.1332 | 21.4 |  | 331
408 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.3 | 1.3 | 1.02 | 17.48 | 3340000000 | 5.0382 | 21.4 |  | 332
409 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.2 | 1.34 | 1.04 | 16.74 | 3540000000 | 4.9431 | 21.4 |  | 333
410 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.3 | 1.37 | 1.06 | 16.85 | 2910000000 | 5.0382 | 21.4 |  | 334
411 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.3 | 1.32 | 1.03 | 16.66 | 5430000000 | 5.0382 | 21.4 | Re-tested | 335
412 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.4 | 1.34 | 1.04 | 16.7 | 2630000000 | 5.1332 | 21.4 |  | 336
413 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.5 | 1.35 | 1.05 | 16.36 | 2560000000 | 5.2283 | 21.4 |  | 337
414 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.3 | 1.37 | 1.05 | 16.57 | 3050000000 | 5.0382 | 21.4 |  | 338
415 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.2 | 1.33 | 1.02 | 17.09 | 2700000000 | 4.9431 | 21.4 |  | 339
416 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.4 | 1.32 | 1.03 | 16.58 | 3400000000 | 5.1332 | 21.4 |  | 340
417 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.4 | 1.33 | 1.03 | 16.8 | 4140000000 | 5.1332 | 21.4 |  | 341
418 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.4 | 1.34 | 1.05 | 16.36 | 3480000000 | 5.1332 | 21.4 |  | 342
419 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.3 | 1.29 | 1.01 | 16.8 | 2680000000 | 5.0382 | 21.4 |  | 343
420 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.4 | 1.34 | 1.04 | 16.69 | 2760000000 | 5.1332 | 21.4 |  | 344
421 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.4 | 1.35 | 1.05 | 16.46 | 2960000000 | 5.1332 | 21.4 |  | 345
422 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.5 | 1.36 | 1.05 | 16.57 | 2950000000 | 5.2283 | 21.4 |  | 346
423 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.3 | 1.35 | 1.04 | 16.81 | 3270000000 | 5.0382 | 21.4 |  | 347
424 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.4 | 1.33 | 1.03 | 16.95 | 3300000000 | 5.1332 | 21.4 |  | 348
425 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.4 | 1.38 | 1.07 | 16.3 | 2750000000 | 5.1332 | 20.8 |  | 349
426 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.4 | 1.35 | 1.04 | 16.78 | 5080000000 | 5.1332 | 20.8 | Re-tested | 350
427 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.4 | 1.34 | 1.04 | 16.83 | 3330000000 | 5.1332 | 20.8 |  | 351
428 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.2 | 1.33 | 1.01 | 17.27 | 2540000000 | 4.9431 | 20.8 | Re-tested | 352
429 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.3 | 1.32 | 1.01 | 17.28 | 2640000000 | 5.0382 | 20.8 |  | 353
430 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.4 | 1.33 | 1.03 | 16.81 | 2530000000 | 5.1332 | 20.8 |  | 354
431 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.4 | 1.35 | 1.04 | 16.86 | 1910000000 | 5.1332 | 20.8 |  | 355
432 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.4 | 1.35 | 1.04 | 16.64 | 2750000000 | 5.1332 | 20.8 |  | 356
433 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.2 | 1.31 | 1.01 | 16.9 | 2620000000 | 4.9431 | 20.8 |  | 357
55 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.3 | 1.29 | 0.99 | 17.68 | 1880000000 | 4.9279 | 14.8 |  | spare
56 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.3 | 1.3 | 0.99 | 17.62 | 2590000000 | 4.9279 | 14.8 |  | spare
58 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.3 | 1.31 | 1 | 17.59 | 2650000000 | 4.9279 | 14.8 |  | spare
60 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.3 | 1.33 | 1.01 | 17.47 | 3070000000 | 4.9279 | 14.8 |  | spare
67 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.3 | 1.33 | 1.02 | 17.5 | 2580000000 | 4.9279 | 14.8 |  | spare
69 | 1087 | XIRE0313MS6 |  | x | 21.3 | 5.3 | 1.33 | 1.01 | 17.49 | 2360000000 | 4.9279 | 14.8 |  | spare
251 | 1087 | XIRE0313MS6 |  | x | 22 | 5.2 | 1.32 | 1.02 | 17.12 | 2730000000 | 4.9296 | 20.3 |  | spare
254 | 1087 | XIRE0313MS6 |  | x | 22 | 5.2 | 1.31 | 1.02 | 16.89 | 2650000000 | 4.9296 | 20.3 |  | spare
288 | 1087 | XIRE0313MS6 |  | x | 22 | 5.2 | 1.36 | 1.04 | 16.86 | 7800000000 | 4.9296 | 20.9 |  | spare
289 | 1087 | XIRE0313MS6 |  | x | 22 | 5.2 | 1.33 | 1.03 | 16.95 | 2700000000 | 4.9296 | 20.9 |  | spare
292 | 1087 | XIRE0313MS6 |  | x | 22 | 5.2 | 1.3 | 1.02 | 16.72 | 2470000000 | 4.9296 | 20.9 |  | spare
298 | 1087 | XIRE0313MS6 |  | x | 22 | 5.2 | 1.31 | 1.02 | 16.68 | 2700000000 | 4.9296 | 20.9 |  | spare
303 | 1087 | XIRE0313MS6 |  | x | 22 | 5.2 | 1.34 | 1.03 | 16.8 | 2910000000 | 4.9296 | 20.9 |  | spare
315 | 1087 | XIRE0313MS6 |  | x | 22 | 5.2 | 1.33 | 1.03 | 16.86 | 2440000000 | 4.9296 | 20.9 |  | spare
362 | 1087 | XIRE0313MS6 |  | x | 22 | 5.2 | 1.3 | 1.02 | 16.84 | 2560000000 | 4.9296 | 21.4 |  | spare
363 | 1087 | XIRE0313MS6 |  | x | 22 | 5.2 | 1.34 | 1.03 | 16.81 | 3320000000 | 4.9296 | 21.4 |  | spare
374 | 1087 | XIRE0313MS6 |  | x | 22 | 5.2 | 1.34 | 1.03 | 17.08 | 2770000000 | 4.9296 | 21.4 |  | spare
383 | 1087 | XIRE0313MS6 |  | x | 22 | 5.2 | 1.31 | 1.01 | 16.93 | 3550000000 | 4.9296 | 21.4 |  | spare
346 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.2 | 1.32 | 1.03 | 16.77 | 2910000000 | 4.9431 | 21.5 |  | spare
348 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.2 | 1.32 | 1.02 | 17.15 | 3010000000 | 4.9431 | 21.5 |  | spare
405 | 1087 | XIRE0313MS6 |  | x | 22.1 | 5.2 | 1.31 | 1.02 | 17.28 | 3180000000 | 4.9431 | 21.4 |  | spare
|  | MIN |  |  | 21.2 | 5.2 | 1.29 | 0.99 | 16.22 | 1850000000 | 4.9279 | 14 |  | 378
|  | AVG |  |  | 21.7521 | 5.3905 | 1.3359 | 1.0287 | 16.9921 | 2877275132.2751 | 5.0752 | 18.2368 | 1.585%
6.07% |  | MAX |  |  | 22.1 | 5.6 | 1.39 | 1.07 | 17.91 | 8810000000 | 5.236 | 22.5
|  | STDEV |  |  | 0.28454 | 0.089587 | 0.020036 | 0.015869 | 0.36204 | 796840400.8123 | 0.080466 | 3.0057
| Max % Less |  |  |  |  | 3.534% | 3.436% | 3.76% | 4.544% | 35.7% | 2.902%
| Max % More |  |  |  |  | 3.887% | 4.05% | 4.017% | 5.402% | 206.2% | 3.168%
[7 empty] | 0.8 | 0.6 | 15 | 0 | 4
[7 empty] | 1 | 0.8 | 15.5 | 2000000000 | 4.2
[7 empty] | 1.2 | 1 | 16 | 4000000000 | 4.4
[7 empty] | 1.4 | 1.2 | 16.5 | 6000000000 | 4.6
[7 empty] | 1.6 | 1.4 | 17 | 8000000000 | 4.8
[7 empty] | 1.8 | 1.6 | 17.5 | 10000000000 | 5
[9 empty] | 18 | 12000000000 | 5.2
[9 empty] | 18.5 | 14000000000 | 5.4
[9 empty] | 19 | 16000000000 | 5.6
5.8
6
(Max-Min)
Average